Galiso Dual Acting Pump

(GDP-10k)

Repair and Maintenance

Component Repair

Parts Explanation

Outlet Check Valve

Isolator Tube

Outlet Check Valve

Isolator Tube

Hydraulic End

Hydraulic End

Inlet Check Valve

Inlet Check Valve

The most common problem in the field with this pump is a drip 			of water from the isolator tube.

The Water drips from here

This indicates that the seals in the hydraulic end have excessive 			wear.

Follow these steps to repair

- First remove both the high pressure (9/16 wrench) and low pressure (7/8 wrench) tubing from fittings located on the check valves.

High pressure sst line

Low Pressure Copper line

- Unscrew the hydraulic end assy from the pump

Hydraulic End

- Set upright to pick out the old seals

Seals will be located here

- Replace existing seals with new and reassemble the pump (when installed the o-ring should be sandwiched between the two Teflon seals)

To repair the check valves of the pump:

1. First remove both the high pressure (9/16 wrench) and low pressure (7/8 wrench) tubing from fittings located on the check valves.

High pressure sst line

Low Pressure Copper line

2. Remove check valve

3 The check valve repair kit consists of the following:

Buna N Oring

SST Check Ball

Spring Retainer

C Style retaining Clip

Spring

Ball Retainer

- Remove C Clip and all Parts.

- Use O-ring pick to remove O-ring

- Replace O-ring, grease O-ring with O2 compatible grease

- Replace all of the hard items and return to service.

To replace the Air seals in the Isolator Tube:

High pressure sst line

Low Pressure Copper line

First remove both the high pressure (9/16 wrench) and low pressure (7/8 wrench) tubing from fittings located on the check valves.

- Unscrew the hydraulic end assy from the pump

Hydraulic End

- Loosen the 1/8” Allen on the end cap assy and unscrew          the Isolator tube

1/8” allen located here

Isolator Tube

4 Remove old seal with o-ring pick

Seals Location

Seal Location

- When installing new seal be sure that the 0-ring is installed into the Teflon backup seal prior to installing in the isolator tube

- Lightly bend in one side of the Teflon Backup/O-ring assy to make a “C” shape and install into the Isolator tube. Be sure that the seals are completely in the groove.