TEST PROCEDURE FOR RECORTEST OPEN SYSTEM

Prior to testing and before the system is turned on, it is important to complete the following.

- ensure the cylinders to be tested are free from defects and have been properly inspected

- completely fill the cylinders with clean water and allow cylinders to achieve room temperature

- master gauge and calibrated cylinder must be on hand to verify calibration each day of testing

SYSTEM START UP PROCEDURE (RECORTEST OPEN)

Read all instructions before attempting to operate the Recortest Open test system. Before turning on

computer, ensure air and water are turned off or closed.

- Turn power to computer, monitor and printer on and wait for computer to initialize.

When system has gone to blue screen, do the following. Go to EDIT in main menu, press enter.

Go to hardware and press enter. Once in hardware screen, press QUIT,(F10) key.

(This must done each time computer is turned on and testing is to be done)

- Turn on air and water to system. Verify that control air pressure 90 PSI on small regulator at right

on pump manifold. Check oil level in lubricator. Place calibrated cylinder in test jacket and connect hoses to test head, blue hose first. Turn 3 way valve to EXPANSION.

- Bleed air from expansion lines by filling bowl using bowl/jacket fill valve then removing blue hose and allowing water to drain back to jacket. Be sure not to drain water below end of probe.

Ensure that 3 way valve is turned to expansion when cylinder is in jacket and hoses are connected.

- Go to VERIFY screen and enter verify for the jacket the calibrated cylinder will be used in

for that day. Ensure that before any checks are done, the system will hold on zero to make

certain the system is stable.

- To verify the bowl reading, place test weights on the bowl hanger (with the reading starting

at 0.0 The reading for the two test weights should be 100 grams or (99.7 – 100.3). If the

reading is outside this tolerance, you must recalibrate the load cell using the calibration

procedure in the manual.

- To verify the pressure, connect the master gauge to the test head and attach the test hose to the

top fitting on the master gauge **. Pressurize the system using the AUTO (F5) key and enter a

pressure that is within the testing range of cylinders to be tested that day. Verify that the master

gauge reads within the 1% required by DOT and that the expansion is within its 1% value.

If the reading is outside the tolerance acceptable by the DOT, the pressure will need to be

recalibrated. Refer to the manual if pressure calibration is needed.

- Bleed off pressure using BLEED (F6) key and verify system returns to 0.0. If system does not

return to zero, ** you must (stretch) once or twice to get calibrated cylinder to zero out.

Calibrated cylinder must show NO permanent expansion.

- Remove master gauge from system and exit verify screen. Go to the test cylinder screen and

test calibrated cylinder within 500 psi of the pressure ranges tested that day. (DOT mandate)

The calibrated cylinder is always the first cylinder to be tested in the day to provide a record

of the system accuracy.

- Once calibrated cylinder has been tested, the operator will proceed to test cylinders that day.

The operator should be familiar with cylinder markings to ensure proper entry into the fields

provided in the test menu. Entries will now be made into the test menu for the next cylinder

to be tested and jackets enabled or disabled.

Prior to cylinder testing, the 3 way valve needs to turned to hold, the hoses removed and the

calibrated cylinder removed and replaced with the first cylinder of the day to be tested. When

that cylinder has been placed in the jacket, fill the jacket, then connect the hoses, then be sure to

turn 3 way valve to expansion before putting pressure on the cylinder. This part of the procedure

is very important to do in the correct order. Any questions, call Galiso.

CYLINDER TESTING

- Using the guides in the REC Open manual the operator will enter the information for each cylinder to be tested and allow the system to automatically test the cylinder. As each cylinder is finished and the results show the cylinder passes, the cylinder may then be removed from the jacket and placed in the inverter, (manual or automatic) to be drained and dried.

- When the days testing is complete, exit the test cylinder menu and go to FILES in the main menu.

To print out the report for the day, go to report and press enter. Continue to press enter until you

see at the top of the screen appears “PRINT THE WHOLE FILE”, press enter. The operator may

choose to select a range of dates or other information to print from. Note- Be sure when printing

results, you verify printout is clear before deleting any information.

- The operator may choose to archive the results to a floppy disk. Follow the directions in the

archive section of printing a report in the REC OPEN manual.

- To shutdown the system, turn off (or close) the water and then the air to the system.

From FILES, go to quit and press enter. The computer may now be turned off.

NOTES- refer to the operation manual (provided) for any detailed information or

call customer service at (800)854-3789 for additional information and assistance.